# MicroEMG Jitter Analysis Summary

Richard Howey, Gabrielle Schroeder

Research Software Engineering, Newcastle University

# Introduction

This document is a brief overview of the jitter analysis implemented by the RSE team in Python using data gathered by the MicroEMG. The analysis consists of 4 basic steps: motor unit identification, fibre potential localisation, fibre potential clustering (identifying fibres) and finally jitter analysis between identified fibres.

# Motor Unit Identification

The (Teager-Kaiser) TK filter [1] was faithfully translated from a MATLAB package. There were, however, some differences:

Rounding. The rounding in MATLAB always rounds up when the last digit is 5. In Python “Banker’s Rounding” is used where if the preceding number is odd the number is rounded up but rounded down otherwise.

Peak finding. The standard peak finding algorithm in Python package scipy.signal has subtle differences to the one in MATLAB.

Arbitrary labelling of motor units. In the original algorithm the motor units are arbitrarily labelled with positive integers 1, 2, 3, etc. In the Python translation the motor units are arbitrary labelled differently. Further processing of the motor units merge ones that are deemed to be the same. Unfortunately, the outcome is dependent on the order on which they are processed.

For the three reasons above the TK filter in Python gives different results from the one in MATLAB despite performing the same algorithm. Points 1 and 2 also give rise to differences elsewhere in the overall algorithm where they are used.

# Fibre Potential Localisation

The fibres were localised by initially reproducing the methods of S. Maitland [2] translated from the original MATLAB code.

The main difference was the 2D peak finding algorithm used within the method. The algorithm in the MATLAB code did not exist in Python so another method had to be used. The 2D peak finding algorithm finds peaks in data from a window of time (x-axis) for every channel/electrode (y-axis). Firstly, the absolute value of the data was taken so that dips are also considered as peaks. A Gaussian filter (standard deviation 2, truncated at 2 standard deviations) was applied to smooth the data. (Note that a top hat transform was not applied but an option is included if desired.) A threshold was used to return all peaks above this threshold. The threshold was adjusted until the number of peaks was between 1 and 5. A neighbourhood size of 5 channels and 10 time-units was used to detect peaks. Equivalent to 500 micro-seconds for data with a sampling rate of 20000 Hz.

# Fibre Potential Clustering

The fibre potential positions were clustered to give estimates of the fibre positions. The algorithm DBSCAN [3] was used to cluster the points as it discards outlying points whereas other clustering methods usually assign every point to a cluster. As the data was quite poor this was a very useful step to improve the quality of the jitter analysis. There are 2 important parameters to the DBSCAN clustering method. Points are defined as core points to a cluster if they contain a certain number of other points within a given radius. The radius was set to 0.1 and the number of points to 10.

The clusters were used to provide an estimate of the fibre positions for the motor unit by using the average position. In plots of fibre positions 2 standard deviations are shown by default to show the uncertainty of their location.

As well as the fibre potential locations, time can be added in as an extra dimension during clustering. This was added and the time values (relative to the motor unit potential onset times) were scaled to be in the interval of size [0, 2.5]. Other scaling options were made available.

# Jitter Analysis

Jitter analysis was performed by following the methods of Dahlbäck [4] to compute the mean consecutive differences (MCDs) between fibre pairs. A heat plot comparing every fibre in a motor unit can then be presented which may highlight if any individual fibre has higher jitter than the others (provided there are at least 3 fibres).

For a given motor unit (action) potential (MUP) there may be 0, 1 or multiple fibre potentials assigned to a fibre (clustered from above step). Ideally there should be 1 for every fibre. Then the MCD can be calculated between two fibres by firstly calculating the time intervals between every fibre potential pair for every MUP. The consecutive differences can then be calculated between adjacent MUPs. The mean then defines the MCD.

When there is more than one fibre potential for a fibre for a MUP then the one with the closet time to the median time across all fibre potentials and MUPs is chosen.

When there are no fibre potentials for a fibre and MUP then the time interval cannot be calculated for the MUP. Therefore, two consecutive differences can also not be calculated. Therefore, if the data is poor there may be few consecutive differences in the final calculation of the MCD. A heat plot can show how much data is used in the MCD calculations.

# Outliers

Outliers are currently discarded in 3 different places in the analysis pipeline.

- Clustering with the DBSCAN algorithm.

In the jitter analysis discarding outliers in time intervals that are greater than 1.5 times the interquartile range (IQR) beyond the IQR boundary.

As 2. above for the consecutive differences.

# Quality Assessment

The are several things that can be done to inspect the quality of the jitter analysis by visually inspecting the analysis output.

The plot showing the clustering can show how well the fibre potentials have been clustered. These should be nice and compact. The 3D plot, with time on the z-axis, can further show how well the clustering has performed. There can be many outlying points in the time axis that can be discard with suitable clustering settings.

The plot of signals for an interval over all MUPs for one channel can reveal if the peaks really belong the same fibre or not. If not, then this will poorly impact jitter analysis.

The histogram of the time intervals (used in the MCD calculation) should show a nice normal distribution type shape. If it has a fat tail or shows a bi-modal shape (more than one peak) then this is indictive of poor data. The standard deviation should also be reasonably small.

Similarly, the histogram of the consecutive differences should also show a nice normal distribution type shape.

The median consecutive differences can also be plotted. If the data is poor, then the medians will be much lower than the reported MCDs. Note that the medians can only take a limited number of values due to the time differences in the signal data. So, if the MCD jitter value is small the median may be 0.

The amount of data used for the jitter analysis can be displayed in a heat plot both as an absolute count and as a percentage of the total possible data.

# Future Work Considerations

Further explore default parameters for entire analysis. It may be possible to improve them while being careful to not overfit them to one dataset.

Review of analyses steps to see if any improvements can be made. The 2D peak finding step is particularly difficult and improvements may be possible.

Identify which parameters are the most sensitive to the final jitter analysis results and identify suitable ranges. These could be made prominent to the user to adjust. However, it is likely that there are suitable fixed values for many parameters.

Set maximum distance in MUPs between differences to calculate a consecutive difference. This would allow more data to be used but would not be strictly consecutive. This parameter is effectively set to 1 currently (if it existed). So maybe not do this.

Improvement of quality control of the data. Currently a simple procedure to discard outliers is used. More sophisticated measures could be explored to clean the data.

# Analysis Parameters

There are many parameters that can be set to do the analysis. Here is a brief list of most of them. The ones that are currently not used or considered effectively fixed are not listed.

Settings for TK Filter options

- tk_filt_thres_spike = 0.1. Threshold to determine a peak.

tk_filt_thres_PsC = 0.1. Pseudo correlation threshold to decide if motor units are similar.

Settings for fibre localisation.

- n_electrodes = 64. The number of channels/electrodes.

- spike_dur = 20. The number of time units around a peak used to calculate the location of a fibre potential.

- half_subsample_size = 200. Half the subsample size to consider when detecting peaks in a MUP.

- xtol = 0.01. Stopping convergence threshold of fibre position used in fibre localization algorithm.

- ftol = 1. Stopping convergence threshold of function value used fibre localization algorithm.

- y_scaling_factor = 2.0. Scaling factor of y-axis for returned estimated positions of fibre potentials.

2D peak finding options for localisation

- find_peaks_2d_min_peaks = 1. The minimum number of peaks to find.

- find_peaks_2d_max_peaks = 5. The maximum number of peaks to find.

- find_peaks_2d_neighbour_chns = 5. Neighbourhood number of channels used to determine a peak.

- find_peaks_2d_neighbour_time = 10. Neighbourhood number of time units used to determine a peak.

DBSPAN clustering options

- dbscan_eps = 0.1. Epsilon value for DBSPAN algorithm. Radius around core point.

dbscan_min_samples = 10. The minimum number of points for a core point to have within the radius above.

# References

[1]	H. Sedghamiz and Daniele Santonocito "Unsupervised Detection and Classification of Motor Unit Action Potentials in Intramuscular Electromyography Signals", in The 5th IEEE International Conference on

E-Health and Bioengineering - EHB 2015,  2015. DOI: 10.1109/EHB.2015.7391510

[2]	S. Maitland, E. Escobedo-Cousin, I. Schofield, A. O’Neill, S. Baker and R. Whittaker, "Electrical cross-sectional imaging of human motor units in vivo", in Clinical Neurophysiology, 2021.  DOI: 10.1016/j.clinph.2021.12.022

[3]	M. Ester, H.-P. Kriegel, J. Sander and X. Xu, "A density-based algorithm for discovering clusters in large spatial databases with noise", KDD'96: Proceedings of the Second International Conference on Knowledge Discovery and Data Mining, 1996: DOI: 10.5555/3001460.3001507

[4]	L. O. Dahlbäck, J. Ekstedt and E. Stålberg, "Ischemic effects on impulse transmission to muscle fibers in man", Electroencephalogr Clin Neurophysiol. 1970: DOI: 10.1016/0013-4694(70)90100-8